## Slide 1: R(T) correction WO3196 Device 9
Rijk Hogenbirk
17/04/2020

## Slide 2: Measurements
Hydrogen to Air measurement
  0324_1904_WO3196dev9_H2ToAir
R(T) fit
  0327_1231_WO3196dev9_Tsteps5C_RvsT

## Slide 3: Dev9 40-100C R(T) fits –only R(T)
0327_1231_WO3196dev9_Tsteps5C_RvsT
Variable | Value | Std | Unit
A | 1.42E+02 | 1.94E+01 | Ohm
Ea | 312 | 3.82 | meV
Fit looks good and Ea is in the same range as the value from Giordano

## Slide 4: H2toAir measurement
0324_1904_WO3196dev9_H2ToAir
Temperature change when pressure is changed causes the recovery time to be impossible to determine. The response to air should be a reverse exponent (), but this is completely obscured by the resistance change due to the temperature changes.
This instability in the temperature is caused by the temperature controller not responding properly to the sudden temperature shift, see mass-spring system response to bump

## Slide 5: Problem
Value of R(T) at 65C is different from the value of R for the H2toAir measurement, so the R(T) curve cannot be used as is to correct the wobble
6.3MOhm (R(T)) vs 2.6 MOhm (H2toAir)
Also the temperature dependent part of the Arrhenius equation () cannot be easily taken out to allow for in in-situ fit, with

## Slide 6: Plan to correct temperature wobble H2toAir measurement
Assume temperature dependence will be the same across measurements, i.e. Ea is constant
Use the free A parameter to adjust Arrhenius equation to the correct resistance range for each point of the H2toAir measurement
Use the Arrhenius equation to artificially correct the temperature to 65C that the device should have the entire measurement

## Slide 7: Correction (1)
Wobble is reduce, but not reduced enough to see the exponential curve underneath
Negative peak at 600s is exaggerated by the correction
Correction seems to be of the right value, so there must be something else happening here
  Look at synchronicity between resistance and temperature peaks?

## Slide 8: Synchronicity
Check time of first negative temperature peak and first negative resistance peak
  Negative peak as resistance is positively correlated with temperature
If there is an offset between the two, implement this in the correction

## Slide 9: Synchronicity first peak
First negative peak at 608.5 s
Double peak with the second one more pronounced
First at 602.75 s
Second at 604.5 s
So there is a time shift between the peaks of at least 4 seconds

## Slide 10: Correction with time shift
Timeshift can only be multiples of the time between measurements
Measurement time (or step) is 0.43 s
Peaks are the smallest for a time shift of 4 steps, but do not disappear enough to show the exponential curve
If the time shift is increase further, the correction overshoots the peaks and creates more and deeper peaks
Dt = 0 steps, 0 s
Dt = 4 steps, 1.72 s
Note: different y-scale for all images
Dt = 20 steps, 8.6 s
Dt = 2 steps, 0.86 s

## Slide 11: FFT Correction
As the temperature wobble is periodic, one should be able to filter it out in Fourier space
To do this, first the FT spectra of the temperature and resistance measurements need to be taken to see which contributions come from the temperature wobble and which from the response to changes in hydrogen
Then some filter needs to be applied to filter out the contributions from the temperature wobble

## Slide 12: FFT Results
Around 50 (Units?) there is a peak in the temperature spectrum and also a slight one in the resistance spectrum
Other contribution to the resistance spectrum is the step function-exponential from the response to air/hydrogen
In the resistance spectrum below 25 there is also some stuff, so this might be the underlying data and the peak around 50 the contribution of the temperature wobble
So a bandpass filter around 50 might filter out the wobble
Resistance
Temperature
Magnitude in this slide means absolute value of the FT, frequency here is the index of the array, not Hz

## Slide 13: FT of response to air
for
As
Real time
Fourier space
Python FFT

## Slide 14: FFT Correction – Subtract Gaussian
Plan:
Subtract Gaussian from temperature spectrum to get rid of wobble
Gaussian (a=55, b=0.0465, c=0.01, ) subtracted from magnitude (parameters chosen such that result is not negative
Does not work that way, use proper filter instead

## Slide 15: FFT – Actual Filter (bandstop)
Bandstop filter between 25 and 65 mHz
4th order Butterworth (scipy.signal.butter)
Filter dip is off from the centre frequency of the peak in the temperature spectrum
Filtering causes weird behaviour in the first part of the measurements

## Slide 16: FFT – Actual Filter (high pass)
High pass filter of the 4th order with f=0.065Hz
High pass filter gets rid of the change in resistance as well as the constant component
This makes it impossible to get out the response time

## Slide 17: FFT – Actual Filter (lowpass)
Lowpass Butterworth filter of the 4th order with cutoff frequency at 0.025 Hz
Leaves step intact, but significantly slows down the response as it takes out high-f terms

## Slide 18: End

## Slide 19: and calculation
Ranges from 53 to 62 Ohm, which is smaller than the A from the Arrhenius fit, but in the same order of magnitude

## Slide 20: Correction (2)
Here, the first wobble is zoomed in on, between 280 and 420s
The correction has a peak before the wobble ha